SG30SS

OPERATING INSTRUCTIONS

(Humidity Option Included)

sg30SS  CONTROLLED ENVIRONMENT CHAMBER

OPERATION, INSTALLATION AND INSTRUCTION MANUAL

RECEIVING YOUR NEW CHAMBER

Congratulations on your recent purchase of a Hoffman Controlled environment Chamber.  When your shipment arrives, please thoroughly examine the shipping crate for any punctures, dents or signs of rough handling.  It is in your best interest to partially remove or open the shipping container in order to examine the unit for concealed damages which may have occurred during shipment.  If the unit is damaged, it must be noted on the delivering carrier’s delivery slip or bill of lading.

GENERAL INFORMATION AND IMPORTANT OPERATING FACTS

This manual has been compiled to aid in the installation, operation and maintenance of your new equipment.  Please take the time to read all of the material in order to become more familiar with your equipment and its operation and to enjoy optimum performance.

UNPACKING YOUR NEW UNIT

The shipping container should remain on your unit as protection against dents or scratches while transporting it to the actual set-up location.  Remove the shipping container only at the last possible moment by following these simple steps:

1.	Using a pry bar or hammer carefully remove plywood crating from the 	top of the crate.

2.	Lift off the top and carefully pry off and remove the sides.

3.	There may be bolts securing the unit to the bottom of the crate.  These 	would be located at each end of the underside of the container.  In order 	to remove these bolts tilt the cabinet up on one end and place a wooden 	block in the center of the container to allow enough area to remove the 	bolts on one end.  Then tilt the container to the other end with the block 	still in the center, allowing enough area to remove the bolts on the 	other end.  Now carefully slide the unit out of the container bottom.

DO NOT LAY YOUR NEW UNIT ON ITS front, OR TOP.

INSTALLATION AND LOCATION

CLEARANCE

Before moving the cabinet to its final point of installation, accurately measure all doorways or passages to assure clearance.

VENTILATION

The final location site of your air cooled unit must be able to provide a large enough quantity of clean air.  The refrigeration system operates most efficiently and trouble-free with cool, dry air circulation.  Avoid locations near heat and direct sunlight where temperatures can be in excess of 100F.  Also, do not select a location in an unheated room or area where temperatures may drop below 55F.  Air supply to the condensing unit is equally important.  Restricting the air supply will place an excessive heat load on the condensing unit and adversely effect its operating efficiency.

IMPORTANT NOTE:  TO ASSURE SUFFICIENT AIR SUPPLY AND CIRCULATION TO THE CONDENSING UNIT, A MINIMUM CLEARANCE OF 12” ABOVE THE CONTROL PANEL AND 3” ON EACH SIDE AND BACK OF THE CABINET MUST BE PROVIDED.  IF NECESSARY, SPECIAL VENTING OR AIR SUPPLY DUCTS MUST BE INSTALLED BY THE INSTALLER FOR THIS PURPOSE.  DO NOT AT ANY TIME OBSTRUCT THE GRILL AREA IN FRONT OF THE CABINET IN ANY WAY, AND NEVER PLACE OR STORE ANYTHING ON TOP OF THE CABINET MACHINE COMPARTMENT.  THESE RULES ARE ESSENTIAL FOR MAXIMUM COOLING CAPACITY AND LONG LIFE OF REFRIGERATION PARTS.

FLOOR LOADS

The floor at the final location site must be level, free of vibration and strong enough to support the total combined weights of your new cabinet plus the maximum product load which might be placed into it.  A fully loaded reach-in model may reach 2000-3000 lbs.  To estimate the possible product load weight, assume that each cubic foot of storage space weighs approximately 35 lbs.  Multiply 35 pounds by the amount of the cubic feet in the cabinet and obtain the product load weight.  For example, a 20 cubic foot refrigerator can hold approximately 700 lbs of product (35x20) and assuming the refrigerator itself weighs 300 lbs, the total combined weight of cabinet and product is approximately 1000 lbs.  Therefore, the floor in this example must be capable of supporting up to 1000 lbs.

Installing Casters and leveling

If your new unit is supplied with swivel casters, they will be packed in the accessory box that came with your cabinet.  Casters should be installed only when the cabinet is close to its final installation site.  to install casters on your new model, place wooden block along the back, at each end.  tilt the cabinet back, using the wood blocks to help hold the cabinet in its tilted position.  locate the large threaded holes on the bottom of the cabinet and screw the threaded caster studs into the mounting holes, closest to the front of the unit.  repeat this procedure by tilting the cabinet in the opposite direction and installing the remaining casters.  make sure the casters are tightened extremely well (see Figure).  if the casters are not installed tightly, the cabinet will be unstable and may sway or rock, which can damage the cabinet.

if the height of a caster needs to be raised, shims must be installed under the casters which need leveling.  extra large washers, available at most hardware or furniture stores, can be used to shim casters.  for maximum stability under extreme conditions, caster plate assemblies are available from the factory.

do not attempt to level casters by unscrewing them from the cabinet and leaving them loose, as this will cause damage to the cabinet and leg hole threads, voiding all warranties.

INSTALLING LEGS AND LEVELING

Your new model is supplied with adjustable type legs for leveling purposes.  Each single and two section model has four leg mounting holes on its case bottom, and three section models have six leg mounting holes.  Legs are packed in the accessory carton from which they must be removed and installed on the cabinet case bottom (see figure 2).  In order to install the legs, carefully tip the cabinet rearward adding four (4) 2” wood blocks underneath and simply screw the threaded leg studs into the case bottom front leg holes.  Repeat this procedure by tilting the cabinet in the opposite direction and install the remaining legs.  Make sure the legs are tightened extremely well or the entire model will sway or rock with each opening or closing of the doors, possibly causing damage to the case bottom.  This procedure should be performed close to the final installation site.

IMPORTANT NOTE:  IT IS EXTREMELY IMPORTANT THAT YOUR NEW UNIT IS PERFECTLY LEVEL FOR PROPER OPERATION.  If it is not level, the following adverse conditions will become apparent:

1.	The doors will not be properly aligned and consequently will not provide a good seal.

2.	Your unit will run excessively due to improper door seal.

3.	An excessive amount of ice will accumulate inside the finned evaporator 	coil.  If allowed to continue, the ice will eventually block the coil and 	the unit will fail.

START-UP PROCEDURE

ELECTRICAL CONNECTIONS

To insure proper operation, your new unit must be connected to a dedicated circuit that can supply the full voltage as stated on the cabinet serial data plate.  For correct voltage, power draw, and wire accommodations, check the data on the serial data plate located on the unit or the tag on the power cord.  Verify that this information exactly matches the electrical characteristics at the installation location.  a surge protector is recommended to protect your equipment.  if the power supply is not steady a brown out protector may be required.

IMPORTANT NOTE:  The condensing unit supplied with all self-contained units is designed to operate with a voltage fluctuation of +10% of the voltage indicated on the cabinet serial data plate.  Full voltage of the correct service, on an individual line not affected by the operation of other electrical appliances, must be available to the condensing unit at all times.  Burnout of the compressor due to exceeding the high or low voltage limits is easily detected and will automatically void the factory warranty.

115V, 60Hz, 1 PHASE CONNECTION

All 115 Volt units are provided with a U.L. approved power cord and polarized plug which is factory installed.

WARNING:  Any alterations to this cord and plug could cause an electrical hazard and will void the factory warranty.  To insure proper operation, this equipment must be plugged into a NEMA 5-15R compatible, grounded receptacle that can supply the full voltage as stated on the serial data plate.

220V, 50Hz, 1 PHASE CONNECTION

All 208-230 volt units are to be permanently connected and are provided with three (3) field wiring leads which exit the electrical console box located in the machine compartment behind the front grill.  This wiring should be connected to the appropriate power source by a qualified electrician and must conform to all local electrical codes.

SPECIAL VOLTAGE CONNECTIONS

When model are ordered from the factory with special, optional voltages, connections should be made as required on the electrical wiring diagram included in the instruction manual.

START-UP CHECKLIST

After your unit has been installed, leveled, cleaned and electrically connected in accordance with this manual, please take the time before start-up to observe the following precautions to assure troublefree operation:

1.	Check that all exposed refrigeration lines are free of severe dents or 	kinks.

2.	Check the condenser fan and evaporator fans for freedom to rotate 	without any obstructions.

3.	Make sure that the cabinet is properly leveled (see “leveling” under 	installation and location section).

The system should run smoothly and quietly in accordance with generally accepted commercial standards  If any unusual noises are heard, turn the unit off immediately and check for any obstructions of the condenser or evaporator fans.  Fan motors, fan blades, or fan housings can be jarred out of position through rough handling in transit or during installation.

CAUTION:  If unit is unplugged or disconnected for any reason, allow several minutes (5-6) before turning the unit back on to allow the system pressures to equalize.  Disregard this procedure could cause an overload and prevent the unit from operating.

OPERATING INSTRUCTIONS

POWER SWITCH TURNS ON/OFF ENTIRE MACHINE

THE H5LA CONTROL TIMER HAS 2 SEPARATE CONTROL CIRCUITS.  THE NO.1 CIRCUIT CONTROLS THE LIGHTS.  THE NO.2 CIRCUIT ALTERNATES DAY AND NIGHT CONTROLS.  PLEASE REFER TO THE PROGRAMMING SHEETS FOR FURTHER INSTRUCTION ON PROGRAMMING THE TIMER.

OMRON E5CC CONTROLS  DAY AND NIGHT TEMPERATURE AND HIGH AND LOW LIMITS.  THE LIMITS MAY BE SET BY DEPRESSING THE DISPLAY KEY.  AL-2 CONTROLS THE HIGH LIMIT, AL-3 CONTROLS THE LOW LIMIT. THESE SHOULD BE SET AT LEAST 5C ABOVE/BELOW THE OPERATING TEMPERATURE (I.E. IF THE LOW OPERATING TEMPERATURE IS SET AT 20C THE LOW LIMIT CONTROLS OF BOTH E5CC SHOULD BE SET AT 15C OR LOWER, IF THE HIGH OPERATING TEMPERATURE IS SET AT 25C THE HIGH LIMIT OF BOTH E5CC SHOULD BE SET AT 30C OR HIGHER).  DEPRESS THE DISPLAY KEY AND THE CURRENT TEMPERATURE WITHIN THE UNIT AND THE TEMPERATURE SETTING WILL BE DISPLAYED.

TO ADJUST OPERATING TEMPERATURE PRESS THE UP OR DOWN KEY ON THE CORESPONDING OMRON E5CC CONTROL UNTIL DESIRED SET POINT IS REACHED.

LOW SAFETY WILL SHUT DOWN THE ENTIRE MACHINE WHEN THE TEMPERATURE FALLS TO SETTING OF THIS CONTROL.  THEREFORE THIS SETTING SHOULD BE SEVERAL DEGREES LOWER THAN THE MINIMUM SETTING OF THE DAY OR NIGHT CONTROLS.  MACHINE WILL RESTART AFTER TEMPERATURE IN CABINET RISES APPROXIMATELY 5 C.

HIGH SAFETY WILL SHUT DOWN THE ENTIRE MACHINE WHEN THE TEMPERATURE RISES TO THE SETTING OF THIS CONTROL.  THIS SETTING SHOULD BE APPROXIMATELY 5 HIGHER THAN THE MAXIMUM SETTING OF THE DAY OR NIGHT CONTROLS.  THE MACHINE WILL RESTART AFTER THE TEMPERATURE IN THE CABINET FALLS APPROXIMATELY 5.

RED LIGHT INDICATES EITHER TOO HIGH OR TOO LOW A TEMPERATURE EXISTS IN THE CABINET.  FIRST CHECK THE SETTINGS OF THE SAFETY CONTROLS WITH A THERMOMETER TO SEE IF THEY ARE SET PROPERLY.  IF THEY ARE, THEN CHECK FOR A POSSIBLE POWER FAILURE OR GERMINATOR MALFUNCTION.  A COMPETENT REFRIGERATION REPAIRMAN SHOULD BE CONTACTED IF REPAIRS ARE NECESSARY.

WHITE LIGHT INDICATES THAT THE MACHINE IS ON THE COOLING CYCLE AND THE COMPRESSOR SHOULD BE OPERATING.  THE CONDENSER FAN RUNS CONTINUOUSLY.

HEATER SWITCH TURNS ON INTERIOR HEATER, WHICH SHOULD ONLY BE USED WHEN AMBIENT OR OUTSIDE TEMPERATURE IS COOLER THAN REQUIRED CHAMBER TEMPERATURE.

HUMIDIDITY CONTROL, [OPTIONAL] CONTROLS THE HUMIDIFIER LOCATED IN THE BOTTOM OF THE CABINET.  THE HUMIDIFIER SHOULD BE PLUGGED INTO THE RECEPTACLE IN THE LOWER LEFT CORNER OF THE BOX.  (DO NOT PLUG ANY OTHER DEVICE IN THIS ]SET THE HUMIDITY CONTROL FOR THE DESIRED HUMIDITY AND FILL THE PAN WITH WATER.  (DO NOT USE DISTILLED WATER IN THE HUMIDIFIER.)  THE DRAIN FROM THE EVAPORATOR COIL ALSO GOES INTO THIS PAN.

LOW PROFILE, CEILING-MOUNT EVAPORATOR ASSEMBLY

The low-profile evaporator system is comprised of a generous sized, evenly matched evaporator and air circulating fans contained within an easily accessible, low silhouette, interior ceiling mounted housing.

After defrost, the fan delay control prevents the evaporator fans from operating until the evaporator coil has reached a temperature of 32F (approximately 20 minutes) thus, minimizing warm air circulation inside the cabinet.  Therefore, during initial start-up, and after each defrost cycle, the fans will not turn on immediately.

SAFETY PRECAUTIONS

The following safety precautions should be followed when operating any appliances:

1.	Always disconnect the power cord before attempting to work on or 	clean any equipment.

2.	Disconnect the power cord when the appliance will be idled for a long 	period of time.

3.	Do NOT attempt to service this unit yourself as removing any covers may 	cause exposure to dangerous voltage.

4.	Your model is equipped with a grounded and polarized plug.  Do not 	defeat the purpose of this plug by removing the ground post or using a 	non-polarized adapter without properly grounding the outlet.

5.	Never connect any appliance to a power source while standing in water 	or with wet hands.

6.	When a replacement part is required, always insist on factory 	authorized 	parts only.

MAINTENANCE

PERIODIC CLEANING PROCEDURE

It is best to clean your unit when the product load is at its lowest level inside the cabinet.  To clean the interior or exterior cabinet surfaces, the following procedure should be followed:

1.	Disconnect your chamber from its power supply and remove all product 	from inside.

2.	Open all doors and allow the cabinet to reach room temperature.  	Remove all shelves from within the chamber.  Wash with a baking soda 	and warm water solution.  Rinse thoroughly with clean water.  Dry 	completely with a soft clean cloth.

3.	Once the cabinet has reached room temperature, wash the entire cabinet 	inside and out with a baking soda and warm water solution.  For slightly 	more difficult cleanups, ammonia or vinegar in warm water can be used.  	Rinse thoroughly with clear water and dry completely with a soft clean 	cloth.  Failure to dry all surfaces completely may cause water stains or 	streaking on the finish.

4.	Return the shelves to their positions and return electric supply power 	to the chamber.

PRECAUTIONS

1.	Never use harsh detergents, cleaners, scouring powders, or chemicals 	when cleaning your unit.

2.	Strong bleaches tend to corrode many materials and should never come 	in contact with stainless steel.

3.	Tincture of iodine, or iron should not come in contact with stainless 	steel.  These solutions, which cause stainless steel to discolor, should 	be rinsed off immediately if contact occurs.

4.	Gritty, hard abrasives will mar the finnish and are not recommended.

GENERAL PREVENTATIVE MAINTENANCE

Performance of the air cooled condensing unit located on the top of your new unit depends exclusively on the amount of air passing through the condenser fins.  Your unit will run more efficiently, consume less energy, and provide a maximum of trouble-free service throughout its lifetime if the condenser is kept clean and an adequate supply of clean air is provided at all times.  Periodically (At least once a month) inspect the condenser coil, which is located directly behind the front top grill, to check for debris or blockage.  If the condenser coil is dirty or blocked, disconnect the power supply to your unit and using a stiff brush, brush the dirt from the condenser fins until the condenser is clear from any debris.  Using a vacuum cleaner with a brush attachment may aid in this cleaning process.  After cleaning, restore electrical service to your unit.

PARTS AND SERVICE

Always provide the cabinet model and serial number (located on the data plate) whenever contacting the factory or your dealer regarding questions or when ordering parts.

PARTS LIST

Timer –H5LA	Compressor –Copeland  033

Control E5CC	Refrigerant R134A 11ounces

Lights: LED – Dimmable (Optional)	Humidity (Optional) – watlow pm6

HOBO Data Logger (Optional)

TROUBLESHOOTING AND SERVICING GUIDE

problem	probable cause	correction

| condensing unit will not start - no hum | 1. line disconnected, switch open. 2.fuse removed or blown. 3.overload protector blown. 4.control “off” due to cold location 5.control stuck in open position. 6.wiring improper or loose. | 1.close start or disconnect switch. 2.replace fuse. 3.determine reason and correct/replace. 4.relocate control. 5.repair or replace control. 6.check wiring against diagram. |
| --- | --- | --- |
| condensing unit will not start - hums but trips on overload protector. | 1.IMPROPERLY wired. 2.low voltage to unit. 3.starting capacitor defective. 4.relay failing to close. 5.compressor motor has a shorted or open winding. 6.internal MECHANICAL trouble in compressor. | 1.check wiring against diagram. 2.determine reason and correct. 3.determine reason and replace. 4.determine reason and replace. 5.replace compressor. 6.replace compressor. |
| condensing unit starts and runs, but short cycles on over-load protector. | 1.additional current passing through overload PROTECTOR. 2.low voltage to unit. 3.overload protector defective. 4.run capacitor defective. 5.excessive discharge pressure. 6.excessive suction pressure. | 1.check wire diagram. check for added components connected to wrong side of overload protector. 2.determine reason and correct. 3.check current, replace protector. 4.determine reason and replace. 5.check VENTILATION, restrictions in cooling medium or refrigeration system. 6.check for misapplication. |
| condensing unit runs, but short cycles on: | 1.overload protector. 2.thermostat. 3.high pressure cut-out due to: (a)insufficient air supply. (b)overcharge. (c)air in system. 4.low pressure cut-out due to: (a) valve leak. (b)undercharge (c)restriction in expansion device. | 1.see problem #3. 2.diffferential must be widened. 3. (a)check air SUPPLY to condenser (b)evacuate and re-charge (c)evacuate and re-charge 4. (a)replace, evacuate and re-charge (b)evacuate and re-charge (c)replace expansion device |
| condensing unit runs, but for prolonged periods or continuous. | 1.shortage of refrigerant. 2.control contacts stuck closed. 3.excessive heat load placed into cabinet. 4.prolonged or too frequent door openings. 5.evaporator coil iced. 6.restriction in refrigeration system. 7.dirty condenser. 8.filter drier CLOGGED | 1.fix leak, evacuate and re-charge. 2.clean contacts or replace control. 3.allow unit sufficient time for removal of latent heat. 4.plan or organize schedule to correct condition. 5.defrost evaporator coil. 6.determine location and remove. 7.clean condenser coil. 8.replace, evacuate and re-charge. |
| start capacitor open, shorted or blown. | 1.relay contact not opening properly. 2.prolonged operation on start cycle: (a)low voltage to unit. (b)improper relay. (c)starting load too high. 3.excessive short cycling. 4.improper capacitor. | 1.clean contacts or replace relay. 2. (a)determine reason and correct. (b)replace with correct relay. (c)correct by using pump down. 3.determine reason for short cycle, (see problem #5) and correct. 4.determine correct size and replace. |
| run capacitor open, shorted or blown. | 1.improper capacitor. 2.excessively high line voltage, over 110% of rated maximum. | 1.check size and replace. 2.determine reason and correct. |
| relay defective or blown out. | 1.incorrect relay. 2.incorrect mounting angle. 3.voltage too low or too high. 4.excessive short cycling. 5.loose or vibrating mounting position. 6.incorrect run capacitor. 7.loose wiring on relay or overload. | 1.check relay and replace. 2.remount relay in correct position. 3.determine reason and correct. 4.determine reason and correct (see problem #5). 5.remount rigidly. 6.replace with proper capacitor 7.tighten all wiring screws. |
| product zone temperature too high. | 1.control setting too high. 2.inadequate air circulation. | 1.reset control. 2.rearrange product load to improve air circulation. |
| suction line frosted or sweating. | 1.overcharge of refrigerant. 2.evaporator fan not running. 3.expansion valve stuck open. 4.expansion valve superheat too low. | 1.evacuate and re-charge. 2.determine reason and correct. 3.clean valve, evacuate and re-charge. 4.adjust superheat to required setting. |
| liquid line frosted, cold or sweating. | 1.restriction in drier strainer. 2.liquid line service valve partiallly closed. | 1.replace drier, evacuate and re-charage. 2.open valve fully or replace if necessary. |
| noisy condensing unit. | 1.loose parts or mountings. 2.tubing rattle or vibration. 3.bent fan blade causing excessive vibration. 4.fan motor bearings worn. | 1.tighten all mounting parts and shroud cover. 2.reform tubing to be free of contact. 3.replace fan blade. 4.replace fan motor. |

SG30SS Controlled Environment Chamber

Limited Warranty

Hoffman Manufacturing, Inc warrants that if your SG30SS Controlled Environment Chamber proves to be defective in workmanship under normal use we will provide, without charge, the parts and labor to remedy any such defect for the period of one year from the date of original purchase.

To obtain service under this warranty you must contact Hoffman Manufacturing, Inc and provide the original serial # and model number.  You may contact Hoffman Manufacturing Inc. by phone 1-800-692-5962 (Int’l. Phone 001-541-327-2242), FAX 1-800-343-6724 (Int’l. FAX 001-541-327-2255), E-mail sales@hoffmanmfg.com.

To receive replacement parts under this warranty the original, defective part must be returned to Hoffman Manufacturing, Inc.  Any postage, insurance and shipping costs incurred in presenting or sending for service are your responsibility.  Packing shall be in conformance with the shipping company’s guidelines and insured if required to cover the cost of the part being returned.

This warranty covers only normal consumer use.  Hoffman Manufacturing, Inc. is not responsible for warranty service performed without prior consent of Hoffman Manufacturing, Inc.  This warranty does not cover products which fail to be properly maintained or fail to function properly as a result of misuse, abuse, improper installation, neglect, improper shipping, damage caused by disasters such as fire, flood, lightening, improper electrical current or service provided without the authorization of Hoffman Manufacturing, Inc.

THIS WARRANTY AND REMEDY PROVIDED ARE EXCLUSIVE AND IN LIEU OF ALL OTHER EXPRESS WARRANTIES AND UNLESS STATED HEREIN, ANY STATEMENTS OR REPRESENTATIONS MADE BY ANY OTHER PERSON OR FIRM ARE VOID.  THE DURATION OF ANY IMPLIED WARRANTIES OF MERCHANTABILITY OR FITNESS FOR A PARTICULAR PURPOSE SHALL BE LIMITED TO THE DURATION OF THE EXPRESS WARRANTY SET FORTH ABOVE.  EXCEPT AS PROVIDED IN THIS WRITTEN WARRANTY, NEITHER HOFFMAN MANUFACTURING, INC. NOR ITS AFFILIATES SHALL BE LIABLE FOR ANY LOSS, INCONVENIENCE, OR DAMAGE, INCLUDING DIRECT, SPECIAL, INCIDENTAL, OR CONSEQUENTIAL DAMAGES, RESULTING FROM THE USE OR INABILITY TO USE THE HOFFMAN PRODUCT, WHETHER RESULTING FROM BREACH OF WARRANTY OR ANY OTHER LEGAL THEORY.

Some states do not allow limitations on how long an implied warranty lasts and some states do not allow the exclusion or limitation of incidental or consequential damages, so the above limitation and exclusion may not apply to you.

This warranty gives you specific legal rights, and you may also have other rights which vary from state to state.

Hoffman Manufacturing, Inc

16541 Green Bridge RD SE

Jefferson, OR 97352 USA

OMRON H5LA TIMER PROGRAMMING INSTRUCTIONS

SETTING THE TIME CLOCK

The clock operates on 24 hour time (0-24 hours).  To start the programming of the H5L timer the clock must be set to the current time and day.  This is done by depressing the mode key and holding it until the words “Time adjustment” appear in the upper left corner.  The day of the week will be flashing and may be advanced by pressing the “plus” key once for each day.  When the correct day is flashing depress the “write” key.  The hour of the day will flass at this time.  The correct hour may be adjusted by depressing the “plus” or “minus” key now.  Once the correct hour is displayed depress the “write” key once more.  The minutes will now appear flashing.  The minutes may be adjusted by depressing the “plus” or “minus” key.  When the correct minutes are displayed depress the “write” key once.  The time of day is now set.

PROGRAMMING

Before actually setting the program the programming steps should be written down and checked as you perform the programming.

Step 1

Depress the model key once to advance to program 1 which controls the lights within the chamber.  “Prog 1” will appear in the upper center of the digital display.  Within Program 1 you may set the lights to come on and go off a total of 8 times.

Step 2

The time set in this step will be the time in which you desire the lights to come on. The hour of the day will be flashing and may be adjusted by depressing the plus or minus key once for each hour.  When the correct hour is flashing depress the write key once.  At this time the minutes will appear flashing.  The minutes may be advanced by depressing the plus or minus key once for each minute of adjustment.  When the correct minutes are flashing depress the write key once.

Step 3

The diagram indicating the status of the contacts will flash in the left lower corner.  If the contacts are open (the bar not touching the two points) the lights will not come on.  If the contacts are closed (the bar is touching the two points) the lights will come on at the time set in program 1.  The contacts may be moved by depressing the plus or minus key once.  If no bar is showing depress the plus or minus key one time and the bar will reappear.  To enter the contact setting depress the write key once when the bar is flashing in the right position.

Repeat steps 2 and 3 to program the lights to go off.

Steps 2 and 3 may be done a total of 8 times if you desire the lights to come on and go off more than once within a 24 hour period.  If you only desire the lights to come on and go off once within a 24 hour period do not program additional  hours, minutes and contacts.

Step 4

Depress the mode key.  The days of the week will appear with one day flashing.  To set the days of operation for Program 1 the day of the week desired must be highlighted.  Use the plus or minus key to turn the highlight on or off.  When the desired setting appears depress the write key once to advance to the next day.

Step 5

Depress the mode key.  “Prog 2” will appear in the upper center of the digital display.  Program 2 is used to control the day and night controls.  Program 2 may be setusing steps 2, 3 and 4 as in Program 1.  When the contacts are closed the controls will operate on the day control.  When the contacts are open the controls will operate on the night control. (If a constant temperature is desired it is not necessary to set Program 2 and the chamber will operate continuously on the night control providing the contacts are set to remain open on the night control and remain closed on the day control).

Deleting from Cyclic Programs (If the “Cycle” key is depressed)

Four steps will be simultaneously deleted from the cyclic program if the program is called up and then the Cycle and Write Keys are pressed in order. The start time display will remain, but the program will be deleted.

Mode Key

LCD

Write Key

Plus Key

Minus Key Key

Cycle Key

DO NOT USE

Manual Override Switch

OMRON E5CC OPERATING INSTRUCTIONS

Nomenclature

PROGRAMMING

Step 1

AL-2 controls the high limit.  AL-3 controls the low limit.  (These should be set at least 5C above/below the operating temperature (i.e. if the low operating temperature is set at 20°C the low limit control should be set at 15C or lower, if the high operating temperature is set at 25C the high limit should be set at 30C or higher).

Depress the mode key.  Continue depressing the mode key until AL-2 displays. This may be adjusted by depressing the UP or DOWN KEYS.  When the desired high limit is displayed depress the MODE KEY one time to move to AL-3.  AL-3 controls the low limit.  This may be adjusted by depressing the UP or DOWN KEYS.  Depress the MODE KEY until  the current temperature within the unit and the temperature set point will be displayed.

Step 2

Day and Night temperatures, and other options can be set by depressing the LEVEL KEY for less than one second. L.AdJ will be displayed.   Depress the MODE KEY until SP-0 is displayed.  SP-0 is the night temperature.  This may be adjusted by the UP and DOWN KEYS.  Press the MODE KEY one more time to get to SP-1.  SP-1 is the day temperature.  This may be adjusted by the UP and DOWN KEYS.  No further adjustment is necessary.  Depress the LEVEL KEY once more to begin controlling.  Refer to programming instructions for parameter settings.

| OMRON E5CC TEMPERATURE CONTROLLER OPERATION MENUS |  |  |  |
| --- | --- | --- | --- |
| LEVEL | DISPLAY | SETTING | VALUE |
| Operation Level | Ct1 | Heater Current 1 Value Monitor | 0.0 |
| *Hold down | LCR1 | Leakage Current 1 Monitor | (value varies) |
|  | AL-2 | Alarm Value 1 | High Limit |
|  | AL-3 | Alarm Value 2 | Low Limit |
| for at least 1 second |  |  |  |
|  | Ct1 | Heater Current 1 Value Monitor | 0.0 |
|  | Hb1 | Heater Burnout Detection | 0.0 |
|  | LCR1 | Leakage Current 1 Monitor | 0.0 |
| Adjustment Level | HS1 | HS Alarm 1 | 50 |
|  | SP-0 | Set Point 0 | Night |
|  | SP-1 | Set Point 1 | Day |
| *Hold down | iNS | Temperature Input Shift | 0.0 |
|  | iNRT | PV Input | 1.0 |
|  | HYS | Hysteresis | 0.5 |
| for at least 1 second | SPRt | SP Ramp Set Value | OFF |
|  | SPRL | SP Ramp Set Valve | SAME |
|  | iN-T | Input Type | 9 |
|  | d-U | C/F Selection | C |
| Initial Setting Level | SL-H | Set Point Upper Limit | 40 |
|  | SL-L | Set Point Lower Limit | 2 |
| *Hold down | CNtL | Control Mode | oNoF |
|  | S-HC | Control Type | StND |
|  | PtRN | Program Pattern | OFF |
| for at least 3 seconds | oRev | Control Operation | oR-R |
|  | Alt2 | Alarm Type | 8 |
|  | ALH2 | Hysteresis | 0.2 |
|  | Alt3 | Alarm Type | 9 |
|  | ALH3 | Hysteresis | 0.2 |
|  | EV-1 | Event Input | MSOP |
|  | EV-2 | Event Input | StoP |